# Relatorio Tecnico SGMed Inspector

Escopo: 21JN101A

# Resumo Executivo

Mudancas encontradas: 24 | Eventos avaliados: 989

# O Que Mudou

| Parametro | Esquerda | Direita | Categoria | Severidade |
| --- | --- | --- | --- | --- |
| Date / time | 08/04/26 06:38:30 | 09/04/26 07:37:40 | operational | medium |
| Hi limit | 414.66 | 501.84 | metrological | high |
| Lo limit | 81.01 | 107.75 | metrological | high |
| Standard dens override | 858 | 854.2 | metrological | high |
| Flow rate 1 | 90.01 | 119.72 | metrological | critical |
| Flow rate 2 | 107.62 | 142.68 | metrological | critical |
| Flow rate 3 | 126.92 | 170.22 | metrological | critical |
| Flow rate 4 | 154.89 | 206.99 | metrological | critical |
| Flow rate 5 | 186.43 | 252.2 | metrological | critical |
| Flow rate 6 | 226.28 | 294.66 | metrological | critical |
| Flow rate 7 | 271.72 | 326.34 | metrological | critical |
| Flow rate 8 | 316.71 | 392.83 | metrological | critical |
| Flow rate 9 | 376.96 | 456.22 | metrological | critical |
| MF/error 1 | 1.00221 | 1.00175 | metrological | critical |
| MF/error 2 | 1.00205 | 1.00183 | metrological | critical |
| MF/error 3 | 1.00207 | 1.00162 | metrological | critical |
| MF/error 4 | 1.00196 | 1.00146 | metrological | critical |
| MF/error 5 | 1.00177 | 1.00114 | metrological | critical |
| MF/error 6 | 1.00169 | 1.00109 | metrological | critical |
| MF/error 7 | 1.00137 | 1.00097 | metrological | critical |
| MF/error 8 | 1.00142 | 1.00091 | metrological | critical |
| MF/error 9 | 1.0011 | 1.00078 | metrological | critical |
| Override | 858 | 854.2 | metrological | critical |
| Override | 858 | 854.2 | metrological | critical |

# Cromatografia / GC

Componentes com mudança: nenhum componente com mudança

Componentes lidos no arquivo carregado: nenhum componente lido

Perfil esperado: Metano, Etano, Propano, Hexano+, Nitrogênio, CO2

Componentes esperados ainda ausentes: Metano, Etano, Propano, Hexano+, Nitrogênio, CO2

# Leitura dos Eventos

Run 2 Flow rate alarm from Normal to Low Low [80] (486x) - Run 2

Run 2 Flow rate alarm from Low Low to Normal [80] (460x) - Run 2

User administrator (Web: 10.0.1.50) has logged in (14x) - Sem run especifico

User administrator (Web: 10.0.1.50) has been automatically logged off (12x) - Sem run especifico

'Print snapshot report' executed by administrator (2000) (3x) - Sem run especifico

Run 2 Flow rate alarm from Normal to Low Low [80] (435x) - Repeticao em janela curta no run 2

Run 2 Flow rate alarm from Low Low to Normal [80] (410x) - Repeticao em janela curta no run 2

User administrator (Web: 10.0.1.50) has logged in (8x) - Repeticao em janela curta

User administrator (Web: 10.0.1.50) has been automatically logged off (8x) - Repeticao em janela curta

Parameter Product 1 standard density override was changed from 854.4 to 858 by CommDrivers (10000) (1x) - Repeticao em janela curta

# Proximas Acoes

Validar diferencas metrologicas antes de promover nova referencia oficial.

Confirmar se alarmes recorrentes representam degradacao persistente ou janela operacional esperada.